## Slide 1
Load Lock Gun2 Status Update & (Revisionist?) History
Load lock gun worked very well in the Test Cave, with very good photocathode lifetime.
Once installed in CEBAF (7/07) the gun functioned well, but with very poor lifetime.
During the 9 months of Physics we used every opportunity to investigate what changed…
  Vacuum
    Extractor gauge indicated vacuum OK, but what about small leak?
    Leak checking suffered from no HV RGA and noisy Prep RGA
    Hi-potting HV ion pump eventually lead to FE pump
    HV chamber received RGA and new pump => found vacuum OK
    New pump broke during bake out
    Vacuum work has not been straight-forward, however, appears OK
  Laser retroreflection
    All measurements in Test Cave extracted the unused laser light
    This is not the case in CEBAF
    Very difficult to troubleshoot/fix with NEG tube and infrared laser
Primary job for the SAD is to understand and fix LLGUN2 problem.

## Slide 2
6/12 – Virgin superlattice (SLSP-5246-1) + Verdi DC green laser installed
  QE ~1% @780nm and ~15% @ 532nm
  Verdi helps because the beam case be easily seen in a safe way
6/17 – Again, we tried to move the gun to see a laser retroreflection
  Never a clean reflection from wafer, often a reflection from metal surface (NEG tube)
  Single NEG bellows constrains location/angle of HV chamber (and photocathode)
  Best effort alignment => lifetime still poor, with “mock up” of Test Cave
  Lifetime still low in 0.25-0.4 mA range w/ fiber laser
  Did not run w/ DC? Decided not right direction…
6/19 – Replace NEG tube with vacuum windows
  Gun side
  Take advantage of venting HV chamber, so replaced ion pump => goes well
  Able to visually inspect photocathode in/with electrodes => looks OK
  Added window
  Valve side
  Take opportunity to add “diagnostics cross” w/ RGA => leaky issues, but finally OK
  Added window
  Diagnosis
  RGA leak check => OK, no leaks found
  Visual alignment => clearly see laser window aperture is OK, although window hazy
  Laser alignment => first time to see reflected Verdi from photocathode => ~1 deg off

## Slide 3
7/2 – Bake HV chamber so we can transfer pucks
Took opportunity to test integrity of HV support tube => OK, laser spot steady
HV chamber bakeout => goes very well
7/7 – Tested puck4 (bulk anodized wafer)
Very reproducible, laser spot reproduces
7/8 – Tested all 4 pucks
Put each in at 6 orientations, measured reflection at 3 locations

## Slide 4
Puck 4

## Slide 5
Puck 3

## Slide 6
Puck 8
Test => cracked

## Slide 7
Puck 6

## Slide 8
Puck | Wafer | Status | Puck | Retro | Mean [deg] | RMS [deg] | Tilt [deg]
#4 | SL5246-1 | Cracked | OK | BAD | -0.55, -0.91 | 1.62, 1.67 | 2.0
#3 | AXT-5mm | OK | OK | GOOD | -0.73, -0.85 | 0.74, 0.86 | 0.2
#8 | S6967-1 | Cracked | No ear | BAD | -0.73, -1.02 | 1.23, 1.51 | 1.4
#6 | SL5245-3 | OK (edge) | Crown | GOOD | -0.61, -0.76 | 0.68, 0.90 | 0.6
Retroreflection Summary
7/10 - To validate our conclusions, we’ll replace two BAD pucks with test pucks:
  One polished to a mirror surface (removes uncertainty of wafer mounting)
  One bulk mounted per “best practice”
  Use “new” QA process (measure retroreflection from rotated puck or sample):
    Mirror reference ~0.05 deg
    Polished puck
    w/ lint free ~0.2 deg
    wo/ lint free ~0.1 deg
    Puck w/ wafer wo/ lint free ~0.2 deg

## Slide 9
7/17 - What’s now & next (last?)…
  Complete puck transfer
  Verify alignment with QA’d pucks (mirror puck & wafer puck)
  If OK, set gun into final position
  Survey group measure position of gun & laser reference line
  Install new 2-bellows NEG tube (vent cross + HV + new pedestal)
  HV + NEG bakeout
  Survey group set solenoid on laser reference line, verify HV chamber still OK
  Recheck laser retroreflection vs. rotation
  If possible, make a photocathode, beam and check lifetime
  If OK, load 4 final pucks
  Pass out